## Feuil1
Identification (from doc manual) | read/write |  |  | value, reading
|  | address (hex) | address (dec) | int16 | uint16 | hex
measured value | r | 0x0000 | 0 | 1 | 1 | 0x0001 | Response ok
decimal place of the measured value | r | 0x0001 | 1 | 2 | 2 | 0x0002 | Response ok
temperature value | r | 0x0002 | 2 | 172 | 172 | 0x00AC | Response ok
decimal place of temperature value | r | 0x0003 | 3 | 1 | 1 | 0x0001 | Response ok
temperature value x10 | r | 0x0100 | 256 | 172 | 172 | 0x00AC | Response ok
DO value (mg/L) x100 | r | 0x0101 | 257 | 4 | 4 | 0x0004 | Response ok
DO saturation value x100 | r | 0x0102 | 258 | 1 | 1 | 0x0001 | Response ok | this register cannot be read in sequence with the previous two registers
Temp calibration | r/w | 0x1000 | 4096 | 6 | 6 | 0x0006 | Response ok | to be done first, write data is tha ref temp x10, read data is the calib offset x10
ODO zero cal, calib in anaerobis water | r/w | 0x1001 | 4097 | 739 | 739 | 0x02E3 | Response ok | to be done first, write data is 0, read data is zero offset
? | ? | 0x1002 | 4098 | 739 | 739 | 0x02E3 | Response ok | calculation coeff for calibration factor 1 (mail Seeed)
ODO slope calibration, calibration in air saturated water | r/w | 0x1003 | 4099 | 2206 | 2206 | 0x089E | Response ok | write data is 0, read data is slope value x100
? | ? | 0x1004 | 4100 | 2206 | 2206 | 0x089E | Response ok | calculated coefficient for calibration factor 2 (mail Seeed)
sensor ID number | r/w | 0x2000 | 8192 | 55 | 55 | 0x0037 | Response ok
? |  | 0x2001 | 8193 | 5 | 5 | 0x0005 | Response ok | device identification ( mail seeed)
? |  | 0x2002 | 8194 | 55 | 55 | 0x0037 | Response ok | device address (mail Seeed)
baud rate |  | 0x2003 | 8195 | 1 | 1 | 0x0001 | Response ok | Write 0 for 4800, 1 for 9600, 2 for 19200
? |  | 0x2004 | 8196 | 2024 | 2024 | 0x07E8 | Response ok | year of factory shipment (mail Seeed)
? |  | 0x2005 | 8197 | 304 | 304 | 0x0130 | Response ok | date of factory shipment (mail Seeed) Day month ?
resst to factory settings | w | 0x2020 | 8224 | NC | NC | NC | write only | Write 0 to reset the sensor and calibrate again
Total addresses that responded Ok to the scanning done with mbpoll software on linux
8224