## Feuil1
Identification (from doc manual) | read/write |  |  | value, reading
|  | address (hex) | address (dec) | int16 | uint16 | hex
measured value | r | 0x0000 | 0 | 1 | 1 | 0x0001 | Response ok
decimal place of the measured value | r | 0x0001 | 1 | 2 | 2 | 0x0002 | Response ok
temperature value | r | 0x0002 | 2 | 172 | 172 | 0x00AC | Response ok
decimal place of temperature value | r | 0x0003 | 3 | 1 | 1 | 0x0001 | Response ok
temperature value x10 | r | 0x0100 | 256 | 172 | 172 | 0x00AC | Response ok
DO value (mg/L) x100 | r | 0x0101 | 257 | 4 | 4 | 0x0004 | Response ok
DO saturation value x100 | r | 0x0102 | 258 | 1 | 1 | 0x0001 | Response ok | this register cannot be read in sequence with the previous two registers
Temp calibration | r/w | 0x1000 | 4096 | 6 | 6 | 0x0006 | Response ok | to be done first, write data is tha ref temp x10, read data is the calib offset x10
ODO zero cal, calib in anaerobis water | r/w | 0x1001 | 4097 | 739 | 739 | 0x02E3 | Response ok | to be done first, write data is 0, read data is zero offset
? | ? | 0x1002 | 4098 | 739 | 739 | 0x02E3 | Response ok | calculation coeff for calibration factor 1 (mail Seeed)
ODO slope calibration, calibration in air saturated water | r/w | 0x1003 | 4099 | 2206 | 2206 | 0x089E | Response ok | write data is 0, read data is slope value x100
? | ? | 0x1004 | 4100 | 2206 | 2206 | 0x089E | Response ok | calculated coefficient for calibration factor 2 (mail Seeed)
sensor ID number | r/w | 0x2000 | 8192 | 55 | 55 | 0x0037 | Response ok
? |  | 0x2001 | 8193 | 5 | 5 | 0x0005 | Response ok | device identification ( mail seeed)
? |  | 0x2002 | 8194 | 55 | 55 | 0x0037 | Response ok | device address (mail Seeed)
baud rate |  | 0x2003 | 8195 | 1 | 1 | 0x0001 | Response ok | Write 0 for 4800, 1 for 9600, 2 for 19200
? |  | 0x2004 | 8196 | 2024 | 2024 | 0x07E8 | Response ok | year of factory shipment (mail Seeed)
? |  | 0x2005 | 8197 | 304 | 304 | 0x0130 | Response ok | date of factory shipment (mail Seeed) Day month ? Day of year ?
resst to factory settings | w | 0x2020 | 8224 | NC | NC | NC | write only | Write 0 to reset the sensor and calibrate again
Total addresses that responded Ok to the scanning done with mbpoll software on linux
Do not change values in addresses in red.
this may cause the device to become damaged or inoperable. (SEEED mail)